## Sheet1
Tobias Kind (CC-by) free to use by attribution
The mass of the electron must be included into accurate mass calibrations (as mentioned in 1983 already)
http://fiehnlab.ucdavis.edu/projects/Seven_Golden_Rules/Accurate_Mass/
ppm Calculator
mass | mass+ e(-) | ppm |  | Experimental M here | 787.425
100 | 100.0005 | 5.4858 |  | Calculated M here | 787.4244
200 | 200.0005 | 2.7429 |  | Result [ppm] | 0.74928
300 | 300.0005 | 1.8286
400 | 400.0005 | 1.3714
500 | 500.0005 | 1.0972
600 | 600.0005 | 0.9143
700 | 700.0005 | 0.78369
800 | 800.0005 | 0.68572
900 | 900.0005 | 0.60953
1000 | 1000.0005 | 0.54858
Code: Wednesday - WOC am Time Slot/Poster Number: 08:30 am Session: New Developments in Hybrid MS/MS Instruments
Parts-per-Billion Mass Accuracy for Proteomic Studies: Toward Zero Mass Error with Broadband FT-ICR MS Analysis
Jeremiah Tipton ; Tanner M. Schaub; Chris Hendrickson; Alan G. Marshall National High Magnetic Field Laboratory, Tallahassee, FL
Novel Aspect: First detailed analysis of proteomic data from a 14.5 T hybrid FT-ICR MS
Introduction: High molecular mass accuracy (HMMA) affords improved confidence in mass spectral assignments. For proteolytic peptides commonly generated from gel-based and shotgun proteomic experiments, MMA limits have been defined for elemental, isobaric, and isomeric composition assignments.[1,2,3] Furthermore, the use of HMMA proteomic data and the effects on informatics have also been described.[4,5,6] This study presents results and statistical analysis of proteomic data collected from a 14.5 T hybrid LTQ FT-ICR mass spectrometer. Measured rms peptide mass errors are 300 ppb: the most accurate broadband peptide mass measurements to date. The impact of ppb MMA on protein informatics (e.g., computation speed and false positive rate) is presented and discussed.
Methods: Biological extracts, followed by gel-based or shotgun proteomic protocols, were analyzed in broadband mode (external calibration) with a modified hybrid linear ion trap FT-ICR MS equipped with a 14.5 T magnet. Spectra were collected in top-3 or top-5 data dependent mode, with precursor ions analyzed by ICR (200,000 resolving power at m/z 400) and MS/MS spectra collected in the LTQ. Data was extracted with in-house software, followed by peptide/protein identification with MASCOT, OMSSA, and X!Tandem. False positive rates were determined by both a target/decoy database [6] and ProteinProphet.[7]
Preliminary Results: Preliminary data indicates a standard deviation of 300 ppb for 2300 peptides (505 proteins) identified from 39 in-gel trypsin digested samples. False positives from a reverse database, searched against MASCOT, revealed only 13 peptides with MASCOT scores above 40, thus we estimate the error rate to be less than 1 %. Also, the fragmentation spectra for the false positives were poor. To further enhance the approach, the statistics for more complex systems, such as yeast extracts or sub-cellular human samples, will be presented. Also, the experimental data will be systematically adjusted to investigate the effect of higher and lower mass error on false positive rate and computation speed.
http://dx.doi.org/10.1016/0020-7381(83)85063-3
Computer processing of mass-spectral data. II. Assignment of formulas to experimental masses illustrating program FZM; effects of electron mass and functional form of exponential scan law on accuracy of mass measurement*1
S. E. Scheppele |  |  | International Journal or Mass Spectrometry and Ion Physics
and Q. G. Grindstaff |  |  | Volume 49, Issue 2, 1 April 1983, Pages 179-209
R. D. Grigsby1 and S. R. McDonald
C. S. Hwang
Abstract
The use of our computer method for the assignment of formulas and specific Z-values to masses acquired at high resolution is illustrated in terms of representative data acquired in mass-spectrometric analyses of a hydrotreated tar-sands oil and a coal liquid. The former results are in excellent agreement with those obtained from the correlation of gel-permeation chromatographic elution volumes with molecular structure. The latter results illustrate the output from program FZM, show the procedure for interpreting high-resolution data for fossil fuels, and produce an elemental analysis in excellent agreement with the experimental result. The use of neutral masses is a standard procedure for scan-law calibration. In this regard, neglect of the mass of the electron in determining the parameters in an ideal exponential scan law produces significant errors in the determination of neutral masses. Analytical consideration of the perturbation in mass measurement thus induced results in a rate constant expressed as a power series in time. Expressing the rate constant as an eighth-order polynomial produces insignificant mass-measurement errors over a mass range of two decades. Time moments were acquired for ions in the 70 eV spectra of perfluorokerosene (PFK) and perchlorobutadiene (PCB) at a dynamic resolution of
24000. The parameters in the polynomial scan law evaluated for each scan by linear least-squares analysis of both the ionic and neutral PFK reference-mass data produce virtually identical errors in the calculated ionic and neutral PFK masses over the mass range 51–269, with the maximum deviation being 5.01 × 10−4 u. The zero-order term accounts for 96% of the rate constant at mass 269, and 88% at mass 51. Thus, the deviation of the scan law from exponentiality increases with increasing time and is handled empirically by a rate constant expressed as a polynomial in time. Use of the resulting polynomials to process the PCB time moments produces single-scan r.m.s. errors of
2.6 p.p.m. and an r.m.s. error of 1.7 p.p.m. for the average of ten scans for both neutral and ionic PCB masses, as expected. Systematic and random errors are approximately equal. The use of neutral rather than ionic masses in scan-law calibration introduces insignificant errors into mass measurement, provided that the scan law has the proper functional form.
Anal. Chem., 77 (10), 3394 -3400, 2005. 10.1021/ac0485942 S0003-2700(04)08594-4
Web Release Date: April 15, 2005
Copyright © 2005 American Chemical Society
Measuring the Mass of an Electron by LC/TOF-MS: A Study of "Twin Ions"
Imma Ferrer* and E. Michael Thurman
http://dx.doi.org/10.1021/ac0485942
http://doi.wiley.com/10.1002/rcm.3102
Importance of the electron mass in the calculations of exact mass by time-of-flight mass spectrometry
Imma Ferrer *, E. Michael Thurman